Biocompatibility of Materials

January 23, 2007

Silicone Breast Implants

Estimated 1-2 Million women in the U.S. have had breast implants (1963-1998).

90% Silicone Gel (about 10% are coated with polyurethane foam.)

10% saline solution – 9% NaCl (still available without restriction).

Reasons for implantation

- Cosmetic enlargement or reshaping

- Reconstruction following mastectomy

Sheath can be made of poly dimethyl siloxane silicone (MW ~100,000). inside is low molecular weight silicone that provides the right feel (mw very low). Smooth surface and “hard” surface

Water does not have proper feel (uses same envelope).

1951 Dr. Pagman

Polyvinyl sponge (see figure)

Fibroblasts grew into the sponge and became painful

1962 Dr. Cronin/Gerow

Silicone (see figure)

Jan 6, 1992

Silicone was banned by FDA.

2006 – FDA reinstated sale of silicone devices.

1976 Medical Device Safety Act

FDA was given legal responsibility to regular surgical devices/implants

Grandfather clause enacted for then-current devices.

1992 medical act was amended to add more power to the regulation.

USA did not require a law where a national implant database is maintained. 1992 this was changed for class 3 implants (most severe). Silicone implants are class3. class3 are devices that are actually penetrating the body.

1950s Japanese women injected low MW silicone gel injected into their breasts. Thousands of women developed necrotic breasts.

Smooth surface implants allowed silicone fluid to diffuse out of the bag.

Freely injected fluid

Possible carcinomas

Cancer Has been reported after silicone injection

Deaths have occurred – severe toxicity

Silicone Breast implant complaints:

Reports of adverse effects (can show after a year or two)

- Implant became hard and painful

  - Fibrous encapsulation – eventually capsule contractions/hardens causing pain

  - Capsular contracture – massaging was suggested to reorient the developing fibroblasts

  - Interference with mammograms

  - Silicone bleed

  - Capsule rupture

  - Formation of calcium deposits

  - Calcification

  - Implant shifting position

  - Polyurethane degradation (rough surface implants)

“All implants bleed Silicone gel through their outer envelope” FDA 1992

Countermeasures to prevent capsule formation

- Drains for hematoma and fluid accumulation

- Antibiotics

- Capsulotomy (exercise)

- Retromuscular positioning

- Steroids

- Polyurethane coating

Second lumen to prevent low MW silicone from passing through

Uses fluorine – high electronegativity, large atomic radius => steric hindrance

Silicones:

1940s F.S. Kipping, University College, England, WAS

Preparing small molecules for optical rotation studies, but was troubled by oils and “gunk”  in his reaction flask (could polarize light)

1943 commercial production, commercial names silicone, silastic.

Of patients that require resurgery (independent experiments)

Exp 1 57% fibrous encapsulation

Macrophages attack surface of material – o2 radicals, peroxides, etc. signaling cascade results in fibroblast action after inability to digest invader.

Exp 2 40% due to gel bleeding

Exp 3 15% calcification – polyurethane (rough surface) seemed to precipitate the formation of calcium phosphate.

Poly urethane problems (cleaves at ester group)

Degradation – hydrolysis

By products eg 2,4 TDA, polymer fragments

Inflammatory reaction

24 toluene Diamine (TDA) was formed during enzymatic attack of polyurethane foam for first 4 days 9and goes away) –paper done

Counterpaper (1994) Luu and White

Found hydrolysis of polyester urethane foam in phosphate butter, ph 7.4

Rate of TDA degradatin did not reach steady state equilibrium until after 120 days+

June 1988 breast implants designated as class 3 medical devices

April 1991 fda requires safety and effectiveness data

Jan 1993 voluntary moratorium on sale of silicone implants

1996 sale prohibited (?)

As a result – companies stopped providing materials, biomaterials field shrank, until law was passed to forbid filing of suits to material producing companies. Biomaterials Access Assurance Act of 1997

Studies have found that there is no correlation between silicone implants and rates of cancer. A number of surveys from reputable institutions have supported this

FDA draft guidance document (for testing)

Chemistry (xlinking, heavy metals, saline filler, etc)

Mechanical (fatigue, rupture, etc)

January 29, 2007

Total Hip implant

Patients develop problems because they have been bedridden

Inveted by dr. John Charnley (british)

Teflon was recommended

Polytetrafluoroethylene = low coefficient of friction

First implants (hundreds) put in had to be taken off.

PTFE creeps under load – deforms. Uneven shape increases wear, flaking off pieces which cause pain.

Material choice in US for last 35 years UHMWPE

Stainless steel femoral stem

Installation

Hammer stem into bone – groves in the shaft lock into place on the bone, goes into marrow

Adhesive – grouting material PMMA (Polymethylmethacrylaat) (plexiglass/Lucite)

Polymerization is exothermic reaction temp may rise to about 45-50c

Cartilage disintegrates at around 42c

Bone necroses at 50c

Still being used today

Making the acetabular cup

Powder – put into mold and compress with higher temp.

Under an electron microscope – you see grain boundaries.

Temperature and pressure were not high enough to create large/single crystals

Powder – extruding machine. Extrude a tube, and then machine out the cup

Modularity – different sized heads and stems depending on nature of patient.

About one million steps per year for average person – one million articulations of the implant per year leading to abrasion.

Barium sulfate ring is placed in the acetabular cup – shows up on xrays and allows for examination of cup position

Securing the femoral stem (continued)

More current procedure – deposit a microscopic porous coating on the surface of the shaft. Osteoblasts have a preference for pores of a diameter between 100-300 micronsb. Causes mechanical stabilization of the stem with natural bone.

Ti 6Al 4V commonly used in orthopedic implants

Plasma spray procedure is used to apply coating to the stem.

5 parameters can be used to control pore size.

Problems:

Flakes of the PE cup can cause serious problems – can migrate into lymph node etc.

Corrosion of the metal itself:

Stainless steel can corrode

Ti 6al 4V is corrosion resistant. (see diagram) – W/exposure to air, TiO2 is formed – ceramic material. Used in shaft

Corrosion w/SS Ni -> Ni++ oxidation. Releases atomic constituents that make up the product. The metal ions interact with surrounding tissue forming metallic-organic compounds.

Nickel sensitivity: nickel is the most prevalent cause of sensitivity in patients. Many patients have pre-existing nickel sensitivity.

Fatigue fracture of the implant.

Stress is placed on the stem w/each step.

Elastic Modulus of materials

Bone: cortex 2-3 e6 psi

Stainless steel 25 e6

Titanium alloys 12e6

Bioglass 4585 5e6

PE (high density) 85-160 e3

Mismatch of elastic moduli

Implant is much stiffer tan the bone itself

When bone is overstressed and understressed, bone is resorbed (osteoclasts). However at proper stress levels bone reproduces.

Replacing polymer with ceramic

Aluminum oxide Al2O3

Less deformation/flaking, but more brittle. No corrosion, less wear

European implants (head/cup, stem) tend to be ceramic either ZrO2 or Al2O3

United states implants – 90% are metal stem, UHMWPE head. Ceramics are FDA approved.

Concern about fracture and failure due to brittleness.

In USA (DR. Robert P Heaney)

300,000 broken hips annually (acetabular cup and the femoral bone.)

80,000 are men

1/3 die within one year. (w/o implant?)

Keeping motion of the acetabular cup to a minimum

Square shaped cup

Screw threads

Hip Joint systems. Which system gives least wear

- M on M

  - Used to be popular, sometimes encounter galling – metals can flow and eventually lock together. Improved metallurgy may have overcome this problem

- M on polymer

- C on polymer

- C on C

- C on M

Ceramics are good in compression.

Polymers

- PTFE

- UHMWPE MW=2-5 e6 daltons, n=100,000-200,000

- PMMA

Metal

- SS

- 90%Ti 6%Al 4%V – has been some concern about Vanadium in terms of its toxicity. No confirmed studies about Ti6Al4V, but Nb has been substituted

- Cobalt-chromium-Mo chromium forms chromium oxide which provides corrosion resistance

Ceramics

- Al2O3

- ZrO2

.

Find particles of UHMWPE, PMMA, Metals, bone at the implant site.

Wear: To impair, consume or diminish by use and friction. The potential for wear of medical devices exist when there is relative motion between two solids in contact under load.

Two types of wear:

Adhesive wear: The mating surfaces of the two solids flow plastically under friction. Load forming junctions which break under tangential forces giving rise to wear debris.

Contact occurs only at a few asperities. Diagram

Abrasive wear when a hard body slides over a soft surface (Co-Cr or SS (316L) over PE) A series of grooves or defects are ploughed in the softer material. This is two-body wear.

Particles of wear debris or foreign particles which are trapped between the contacting, sliding surfaces and cause abrasive wear of the surfaces by ploughing action. Also called three body wear. What happens to the material that comes out of the grove?

Femoral surface also undergoes wear, even though it is much harder than the cup.

January 30, 2007

Metal backing on the acetabular cup – prevents distortion and encourages ingrowth (plasma spray)

Fatigue wear: abrasion caused by cyclic loading and lost of material by spalling of surface layers. (spalling = peeling process)

Fretting wear: occurs as a result of frictional oscillatory movement of contacting solids

Abrasive index= Rx/Rs

Rx = test specimen

Rs = standard material

(Measure the number of cycles required to abrade .1 inch of material away.)

What kind of abrasion test should we use?

- pin-ondisc

- ring-ondisc

- journal-and-bush (rod in bushing)

- block on journal (block on rod)

measuring:

- weight loss

- volume loss

- delta thickness/length

- holographic surface measurements

- SEM

- light microscopy

- profilometer (can provide 3dview, valleys and peaks show wear)

TI 6Al 4V is not used in heads on femoral stems because of poor wear characteristics – see oxide needle model.

Particles that flake off are 50-200 um. Macrophages cannot ingest – form giant cells

~10um, macrophages can engulf and digest.

1um = 1 micron = 10e-6 meters = 3.9e-5 inches

10g of tissue containing F75 alloy wear debris could contain 2.9e11 particles w/combined surface area of 1600cm^3.

Dr. Ian Clarke

Clarke hypothesis:

- regardless of

  - implant design

  - material selection

  - fixation mode

relative motion at articulating surfaces or micromotion at any interface may result in the release of microsized particles which provoke significant peri-implant bone loss

- activated macrophages are primary agents in removing debris

where is wear taking place/

-the bone/acrylic cement interface bond may be broken by micromotion leading to production of on a fibrous membrane

further micromotion will produce bone and acrylic debris resulting in macrophage activity, inflammation and bone resorption (leading to three body wear)

when PMMA bone junction is broken due to micromotion – can form fibrous layer and prevents mechanical bonding from taking place

90-95% hip implants in US use UHMWPE

Parameters to consider for UHMWPE

- Oxidation index – important in terms of wear and gamma radiation (used in sterilization). Implant can crumble.

- Molecular weight affects TS, hardness, Elastic mod (2-6million typicall for these products. Too high, gets brittle/stiffer)

- Molecular weight distribution – can purify by dissolving the polymer, heating to 135c. run gel chromatography, or add alcohol and precipitate batches of material at different MW.

- Microstructure – amorphous or crystalline? Desire crystalline. Amorphous glasslike does not have as desired properties

- Is this sample branched or linear? Can affect packing desnsity. LDPE/HDPE

- Radical formation during storage

  - Ethylene oxide sterilization – UHMWPE absorbed ethylene oxide and could diffuse out into the patient

  - Irradiate in nitrogen atmosphere – inert and don’t generate free radicals.

Another case of wear debris: implant loosening

- Micromotion between femoral stem and the acrylic bone cement produces polymer and metal particles

- Transport of these particles can result in accelerated wear of the UHMWPE acetabulum

Osteogenic regulatory molecules

- Bone morphogenetic proteins foster new bone regeneration (BMP-2)

- Sustained release carrier systems

  - Biodegradable collagen scaffolds

rhBMP-2 induces

- Cell migration, proliferation

- Differentiation of mesenchymal cells

- New bone formation

- Increased vascularity

Delivery systems for bone morphogenetic proteins

- Absorbable collagen system (ACS) – BMP-2

- Biodegradable polymer PLA-DX-PEG block copolymer containing BMP-2

Osteoconductive

- Provides a passive structure into which blood vessels may enter and new bone may form

- Graft osteoconduction: the facilitation of blood vessel incursion and new bone formation into defined lattice structure

Osteoinductive

- Contains factors which induce the differentiation of mesenchymal cells into osteoblasts

- Graft osteoinduction: new bone produced the active recruitment of host mesenchymal stem cells from the surrounding tissue, which differentiate into bone forming osteoblasts this process is facilitated by the presence of growth factors within the graft, principally BMPs

Recombinant human bone morphogenetic protein-2 RhBMP-2

Can be produced in cell culture or body

Failure modes:

Cyclic fatigue

Fracture

Surgeon contribution to failure

Poor alignment, selection , etc

The most common cause of failure is aseptic loosening (loosening of stem itself)

Characteristics

- Fibrous membrane

- Wear debris: PMMA, UHMWPE, metal; key factors in implant loosening

- Macrophage activity

- Inflammation

- Bone loss

February 5, 2007

Dental implants

Controversy with Calcium Hydroxyapatite (hydroxylapatite) HA

Testing:

In-Vitro

In-Vivo

Clinical

Ti6Al4V 4 types of screws:

Smooth surface

Roughened surface

Plasma sprayed Ti6Al4V forms porous surface

CaHA surface

Too much hydrogen in Ti tends to embrittle it.

Bone, Dentin, Enamel – all have about 36% Ca, ~17% P. Ca/P = ~1.6

CaHA C = 40%, P = 18.5% Ca/P (ratio) = 1.67

Similar to natural material, hopefully does not elicit response from body.

Plasma spray parameters:

Plasma current

Gas flow rate

Gases

Powder flow rate

Gun distance

Angle of incidence.

Control pore size and density

Want to allow osteoblasts to get into the material.

Coating thickness

Porous Ti6Al4V substrate 60-125 microns

HA 35-50

Total 95-175 microns

What is the crystallinity of the HA

Used x-ray diffraction pattern to find the amount of crystallinity vs amorphous material.

Coating tensile strength was found and determined sufficient

In-vivo testing (dog)

Window cutout in the implant to allow for observation of cell growth

Osteoblasts are sensitive to the microstructure of implant – bare Titanium shows less bone growth (osseointegration) when compared to HA that is more crystalline in features. (65% crystallinity)

Clinical

Base/stem is installed into patient first, (3-months pre cap installation) to allow for the bone integration.

Cap the screw thread hole such that no tissue migrates into the hole, remove when abutment is installed.

Hydroxyapatite (HA) Coatings

What is the controversy?

Want to mechanically stabilize the device ASAP

Calcium Phospates

Ceramics that includes a whole rage of amorphous anc rstalline materials

HA is in this group AKA bone mineral

Whether HA is bioactive and reasons why some CaP are bioactive are not agreed upon. Do they cause more harm or good?

Question of solubility

HA is prone to disappear into solution

TriCalciumPhosphate is highly soluble. TCP Ca3(PO4)2

Response to HA “too soluble” -> fluoroHA

February 6, 2007

A = Machined

B = roughened

C = Ti6Al4V Porous

D = Ti6Al4V Porous substrate, HA porous coated.

Sacrifice time 0, 4, 8, 12, 16 weeks

Osetoconductivity: D>C>B>A

Fibroblasts competed against osteoblasts

In Sample A, Fibroblasts out-competed osteoblasts.

Purchased HA had wide array of purities and xtallinity.

Elastic modulus – want to have similar between bone and implant because of bone growth properties. Want bone to endure stress sucht hat it will not degrade and grow.

The use of polymer Polymethleneoxide was tried in implants (similar elastic mod). Failed because of inadequate adherence.

HA was reported to be soluable in the body – related to the quality of the HA being provided by manufactures.

HAF was suggested as a replacement Ca10(PO4)6F2

Test w/HAF

Samples:

A: surface roughed

B Ti6Al4V porous coated

C: Ti6Al4V + HA porous coated

D: Ti6Al4V +HAF porous coated

Evaluation (Triplicate tests):

Histology of implant interface/contact (% coverage/time)

Mechanical test – force to push implant out of bone.

Affinity index = length of bone in contact with surface / length of implant surface

Summary

- mechanical test

  - samples a/b yielded similar results for 0-12 weeks. Sample b at 18 weeks gave the highest pushout value. Sample a at 18 weeks gave lowest push out value

  - both samples a/b increased in pushout values up to 8 weeks, then declined to minimum at 10 weeks

- histology (see notebook)

histology rating scale

0 0% new bon-implant contact

1 25%

2 50%

3 75%

4 100%

Controversy of HA

Those opposed to HA coating claim:

- it is a brittle ceramic subject to fracture and delamination

  - early ceramic implants were brittle and fragile – new advances have reduced these problems and are not frequent occurrences anymore

- it is easily abraded resulting in wear particles. poor wear qualities.

- it is soluble in vivo leading to dissolution of bond anchoring platform

  - source is the impurity of HA from suppliers.

- the porous coating provides open sites for bacterial infection

- porous coating increases opportunity for metallic corrosion

  - not all the surface is coated by the ceramic material

  - Titainium oxide is very corrosion resistant

- properties of HA coating vary with degree of crystallinity

- plasma spraying of HA chemically transforms material

  - forms metal oxides, carbonates, etc. coating is not pure CaHA

- once fully bone in-grown, it is very difficult to remove the porous implant if it fails in service

Those who support HA coatings claim

- the chemical composition of HA is very similar to hydroxyapatite in bone

- HA has been demonstrated to be osteoconductive

- HA coating promotes more rapid bone in-growth and mechanical stabilization

  - Much improved in the early periods

- when the porous HA coating is deposited on a porous metallic substrate to form a composite material, brittle fracture is minimized or eliminated

  - crack propagation will be arrested and catastrophic failure is less a concern.

- the composite nature of the HA-metallic coating improves abrasion resistance

HA particles are used to repair maxillary and mandibular bone defects

- the solubility of the HA coating can be controlled by controlling the xtallinity

with the porous metal substrate, accidental dissolution of the HA coating does not removed bone anchoring platform

- porous structures do invite potential bacterial infection, but

  - other porous implants are successfully used in the body, EG. goretex, porous ceramics

  - antibiotics may be used to ward off bacterial infection

- porous ceramic HA coatings leave less metal exposed

What is the situation today

- not all HA coatings have the same properties or performance characteristics

- while there was a great increase in HA coatings fro about 1980-1990 there has been a decline in recent years

- there are reports in the literature of

  - failure of HA coated implants

  - success of HA coated implants

- more long term, well designed, statistically controlled implant experiments are needed.

February 12, 2007

Cardiovascular system

Ball and cage:

1952: First heart valve was ball and cage. Dr. Hufnagel

1960: Drs. Harken and Starr placed a heart valve into a patient.

Both had problems with thrombosis and blood coagulation

1969: Occluder disc valve – Bjork-Shiley

1970s Bileaflet – introduced by St. Jude

glutaldehyde xlinks in the leaflets and makes it more durable

Mechanical route -> requires blood thinner for duration of implant.

Chemicals:1) wafarin 2) coumadin

Porcine for older people

Mechanical for younger

Mitral -> LA to LV

Aortic -> LV to aorta

Tricuspid -> RA to RV

Pulmonary -> RV to pulmonary

Valve troubles are usually left side, mitral valve.

Valve replacements typicall last ~20 years now.

Valve functions:

- to deliver the blood with minimal adverse effect

Failures

Ball and cage: ring needs to be welded. Needs to be sewn in place.

Ball swells – lipids are absorbed by the silicone material

Used a metal instead. Stellite 21 Co-Cr-Mo (made noise)

Textile was sewn around the ball to act as a cushion. Textile abrades away, causing clotting

Pyrolytic carbon was a good choice for material – reacted well in blood. Does not clot blood. Good against thrombosis

Hydrodynamics of blood flow is important. Don’t want turbulence or stagnation. Want uniform blood flow to minimize clotting risk.

Occluder valve -> minor orifice and major orifice. Blood flow is impeded by metal wires, also stagnation in the minor portion.

Bileaflet – still major/minor, but no flow obstructions.

Problems resulting from adsorption on synthetic implants

- clotting

- mechanical impairment

- removal of components (important proteins, etc)

Blood clots occur on the sewing ring.

Surgeon can affect performance with his sewing technique -> thread can collect protein.

Fatigue fracture on the occluder strut. 40m cycles / year

Fracture on the pyrolytic carbon disc

Porcine heart valve uses textile sewing ring

Large problem with porcine heart valve -> mineralization, calcification

Sequence of events for intrinsic mineralization

- leaflet tissue is infiltrated by host plasma proteins

- infiltration is a normal event and my provide internal lubricity

- in rare instances host proteins concentrate to the point at which fibrin forms and mineralizes. Tissue surface is still intact

- Extension of mineralized areas disrupts leaflet

Mineralization control measures for synthetic filmes

- Avoid surface imperfections

- Avoid inclusions

- Avoid juvenile subjects – more calcification systems

- Design devices with maximum strain relief

- Use mineralization resistant materials

- Use anti-mineralization chemical treatments.

Crosslinking proteins in the porcine leaflet could possibly be related to calcification

Failure of cardiac valves

Design, engineering: Wear, fracture, poppet escape, cuspal tear, calcification, hemolytic anemia, noise

User: sutures, implantation

Amount of protein absorbed

Extension of proteins on surface (height of polymer/protein above the surface)

Difference in surface energy is attracting other molecules

Imporatance of protein adsorption to synthetic implants

Protein depletion

Adsorption or adherence of physiological components to biomaterials surfaces is a key part of biocompatibility.

Protein adsorption

Albumin -> no platelet adhesion

Nonthrombogenicity

Fibrinogen, gamma globulin, prothrombin -> platelet adhesion, can lead to thrombosis

Artificial Hearts

Barney Clark – first artificial heart. Dr. Robert Jarvik invented the heart.1982.

Biomaterials in the total artificial heart.

Polyurethanes – blood chamber housing, flexing diaphragms, etc

Poly carbonate: valve holders, air chamber base

PVC – drive lines (lots of plasticizer, diffuses out)

Polypropylene – suture material

Woven Dacron – outflow grafts

Darcon Velcro – ventricular anchor

Darcon velour – skin buttons,

Silastic – skin button

Woven silk – aspiration port sealing tie

Thrombus formation – patient surface geometry

Patient thrombus potential, anticoagulation drugs

Surface – surface quality, chemistry

Geometry optimum flow conditions

TAH or VAD?

Christian Bernard 1964 first MD to do a human heart transplant

February 13, 2007

Heart valves

Syntehtic materials

- Polymers, metals/alloys

- Ceramics

Modified Natural materials (tissue valves)

- Bovine pericardium valve

- Porcine valve

Metals:

Stellite 21

Co-Cr-Mo

Ti6Al4V

Stainless steel 316L

Polymers: (see figures)

PTFR

Darcon

Delrin

Ceramics:

Al2O3

ZrO2

Pyrolytic C

Low temperature isotropic ~1000c LTI

Ultra low temperature isotropic ~25c ULTI

General types

Caged-ball

Caged Disk (similar to ball)

Tilting-disk – cage is made of stellite

Tissue

Bi-leaflet tilting-disk

First models: Medtronic (caged), St. Jude (bileaflet), Bjork-Shiley (mitral)

Low profile vs high profile

How far does it reach down into the heart chamber. Want low profile. Depends on shape of material, etc

High profile -> silastic poppett

Problems

Blood clotting

- Occluder opening influences hydrodynamics

- Low or stagnant blood flow leads to thrombosis

- Bjork-shiley 60degree C/C heart valve improved hydrodynamics

Kay-Shhiley heart valve

Caged disk.

Problems: disk gets dislodged, large amount of turbulence, wear

Entire occluder can wear on edges, and lock the valve in one position (partially open, etc)

Delrin was not a satisfactory choice for the occluder

Tissue heart valves

Porcine

Bovine pericardial

Advantages

- Freedom from anticoagulant drug (warfarin/coumadin (rat poison))

- Central orfice flow (hemodynamics)

- Cross-linking stabilization

Disadvantages

- steady degradation with time

- calcification (calcific degeneration)

- inflammatory cell infiltration

- fragmentation of collagen fibers

- leaflet perforation

Types of calcification

Extrinsic

- thrombosis related calcification on surface of leaflets of valve

- hydroxyapatite deposit

intrinsic

- in interior of leaflet along collagen fibrils

See reaction (formation of chelated compound)

Effect on patient w/cardiovascular implant:

- protein deposits (within seconds)

- platelet activation (fibrinogen)

- blood clots

- embolism

- hemolysis (lysis of blood cells)

Affect on implant

Wear

Fatigue/fracture

Blood clotting

Valve malfunction

Hindered hydrodynamics

Dr. Fred Schoen study

In 1980s:

5-year survival 70-80%

10 yr 55-70%

Improved in the 1990s

Jarvik heart patients

- barney clark – 1982 oversized heart removed, lived 112 days

- William Schroeder lived 620 days, strokes, kidney failure, troubled history

Strokes, convulsions, fever (infection)

Biomed, Inc. Danvers, MA 20 years of research -> Abiocor heart

Abiocor Artifical heart (2001)

- Totally implantable

- Permanent

- Rechargeable internal battery

- Powered transcutaneously by external power pack

- Weight: 2pounds

- Size: small grapefruit

- Materials: Titanium pump housing and propeller

- Polymer ventricular chamber

- Constant flowrate

Picking materials

Non-thrombogenic – negative charge on endothelial lining. Repels platelets and blood cells

PTFE has negative charge

Pyrolytic carbon

Blood clotting: Albumin likes hydrophilic surfaces (wetting angle ->small as water spreads across surface)

Want hydrophilic –OH groups, something like PVA

February 19, 2007

Thrombogenicity

Thrombus: blood clots, network of fibrin, platelets, erythrocytes, leucocytes

Thrombosis: formation of a blood clot

Thrombus formation on biomaterials

- Adsorption of plasma proteins

  - Fibrinnogen and thrombin

- adhesion of platelets

- aggregation of platelets

- fibrin formation

- mural thrombosis

endothelial cells typically have negative charge and repel blood cells. When injured, they may have a positive charge and attract platelets, etc to themselves.

Platelets are normally disc shaped. Under activation, they transform to pseudopods (long tubes..?)

| Platelet factors involved in coagulation |  |  |
| --- | --- | --- |
| PF | Name | Action |
| Pf1 | Coagulation factor v | Binds to platelet membrane, can form pro-throbinase -> thrombin |
| Pf2 | Finbrinoplastic substance | Acclerates clotting of fibrinogen via thrombin |
| PF3 | Lipoprotein thrombo-plastin | Phospholipids which provide catalytic surface |
| PF4 | Proteins | Has anti-heparin action |
| Pf5 | Fibrinogen | Attaches to membrane and promotes platelet adhesion |
| PF6 |  | Anti fibrinolytic activity (prevents breaking down fibrin) |
| PF7 | Co thrombolastin | Promotes thrombin formation |
| PF 8 | Von willebrand factor plasma glycoprotein | Promotes platelet adhesion |

Platelets

- contain glycoproteins, GPIIb/IIIa on cell membrane platelet receptors

- platelets must be activated for GPIIb/IIIa to bind to fibrinogen

- cross-linking of GPIIb/IIIa leads to platelet activation

- platelets do not possess a unique adhesive receptor (ie RGD) for albumin

- Platelets generate a chemical factor leading to formation of thrombin, then fibrin

- Activated platelets produce alpha-granule release (pro-coagulant.)

Preparing blood biocompatible surface

- Make it negatively charge

- Coat with albumin

Proteins compete for the surface of an implant

Which protein gets to the surface first makes a difference to the bio-compatibility outcome

Albumin vs fibrogen albumin adsorption decreases platelet adsorption and decreases thrombogenicity

(no receptor on platelet for albumin adhesion)

Albumin does not contain the RGP peptide sequence

Albumin

Adsoption decreases platelet adhesion

Decrease thrombogencity blocks ibrinogen adsoption which attracts platelets

Abumin does not contain RGD (arginin-glycine-aspartic acid)sequence, an amono acid sequence common to adhesive proteins

Albumin must occuy most of the surface (98%) to be anti thrombogenic

Fibrinogen

Adsorption (deposition)

Undergoes conformational change

A condition leading to platelet adsorption

Fibrinogen adsorption increases thrombogenicity by generating fibrin

Fibrinogen

Provides signal to promote platelet adhesion

Thrombin links individual (3) chanins in a bibrinogen single molecule

size

Albumin 69000 MW

Hemoglobin 64450 MW

Globulin 165,000MW

Fibrinogen 400000 MW

Blood clotting factors

I. Fibrinogen

Provides signal to platelets for activation

Reacts with Prothrombin II -> fibrin

II. Prothrombin

Reacts with proaccelerin (V + Ca2+) -> Thrombin IIa

III. Tissue factor

Key to extrinsic pathway

Reactins with proconvertin (VII + Ca2+) to form complex

IV. Ca2+

Acts as catalyst (400x quicker w/Ca2+)

Only non-protein clotting factor

Activates Stuart factor (X-Xa)

V. Procaccelerin

Reacts with prothrombin II -> thrombin IIa

VI. Activated V

Reacts with prothrombin II -> Thrombin IIa

VII. Proconvertin

Reacts with Tissue Factor III + Ca+2 to form complex

VIII. Antihemophilic factor

Activates stuart factor X Ca++ -> Xa

XI. Christmas Factor

activated IXa involved in activation of stuart factor X -> Xa

X. Stuart factor

Involved in conversion of prothrombin II -> Thrombrin IIa

XI. Plasma Prothrombo – plastin antecedent

w/Ca++ activates Christmas factor IX -> IXa

XII. Hageman factor

Activated XIIa is the key to the intrinsic pathway

XIII Fibrin stabilizing facor

Reacts with thrombin IIa for XIIIa

Plays a role in Fibrin formation

Factors leading to the intrinsic pathway

- contact between blood elements and a surface

- damage to the wall of a blood vessel (endothelium)

  - activation of Hageman factor (XII) by

    - exposure of a nonendothelial surface such as collagen (electro0negative and thrombotic

    - platelet membrane-electronegative

    - contact with a foreign substance (implant)

Will a Cl- ion get to the surface first?

Factors leading to the extrinsic pathway

- release of tissue thromboplastin (tissue factor III) from cells external to the vascular processes. It is released when the tissue is damaged

  - in conjuction with proconvertin (factor VII), and Ca2+ it activates the Stuart factor (X).

February 20, 2007

Blood Flow rate

Low Shear thrombosis (venous wounds)

Red thrombosis (erythrocytes + fibrin)

High shear thrombosis (arterial wounds)

Shear stress > 3000 Dynes/cm^2

White thrombosis (Platelets + Fibrin)

Vascular surface

- lined with endothelial cells

- subendothelium layer containing elastin (cross-linked polypeptides)

(factor III usually is related to events outside the vascular system)

Blood Compatiblity: Clinical manifestations

- Small diameter vascular grafts fail early due to thrombotic occlusion

- Synthetic venous prostheses do not exist

- Embolic complications are noted with artificial hearts

- Embolic problems are frequently observed with catheters

- Non-tissue heart valves require lifelong anticoagulation

- Sensors “foul” due to thrombus formation

- Long term implants are seen to be continuously platelet consumptive

- Significant blood damage is observed during hemodialysis and extracorporeal oxygenation (also in heart valves)

Blood coagulation and electrostatic repulsion

- RBCs and platelets have a negative charge

Heparin sulfated molecule discovered at Hopkins early 20th century (Dr. Vincent Gott 1963)

Sulfated muco poly saccharide

Contains –SO3(-1) groups -> key to the anti-thrombogenicity

Found to interfere with factor XII

Factors for biocompatibility

- mechanical

  - tensile strength

  - elongation

  - elastic modulus

  - compressive strength

  - fatigue cycling

  - fracture toughness

- Physical factors

  - Size

  - Shape

  - Sharpness of corners (want round)

- Electrical properties

  - Surface charge

  - Dipole moment

- chemical factors

  - composition

  - surface – hydrophilic/phobic, smooth/porous

  - absorption of H2O, lipids. ElasticM goes down

  - leaching -> rigid, stiffer, EM goes up

  - oxidation

  - xlink

effect of environment on the mechanical properties of materials (ceramic)

| Condition | Crack velocity (cm/sec) |
| --- | --- |
| Dried N2 | 5e-6 |
| .001% h2 | e-5 |
| 50% humidity | e-2 |

Microcracks are very susceptible to moisture

Physiological factors

- biological processes

  - activation of osteoclasts/blasts

  - coagulation cascade

  - protein absorption

  - monocytes -> macrophages -> giant cells

  - skin allergy

- chemical reactivity

  - BMP

- nature of tissue (implant site)

  - hard (osteointegration)

  - soft tissue (fibroblast encapsulation)

- Immune system

Polymers

Change in MW, MWdistribution

Degradation products of materials

Either mechanical or chemically

Calcification

Physical

- change in size

- shape

- surface topology

- optical properties (contact lens – protein adsorption)

- change in hydrophilicy/phobicity

Implant lifespan depends on:

- the surgeon

- patient

- implant site

- implant design

- implant material

- fabrication and processing conditions

- conditions of use

Sulzer swiss medical device company. Ceramic hip implants

2006 experienced many failures due to improper cleaning/fabrication

February 26, 2007

Effect of the implant on the body itself

Wound healing

Fibroblast

Collagen

Macrophage

Potential for infection

Biological

- bacterial infection

- macrophage action

- tissue ingrowth

  - mechanical stabilization

- tissue-implant bond

Bioglass fortified 45S5 (SiO2, p2O5 63%, CaO 34.5 %, Na2O)

Within phase diagram: Good biocompatibility, good cell growth.

Modes of biomaterials degradation

| Material | Chemical | mechanical |
| --- | --- | --- |
| Polymers | Oxidation Hydrolysis Leaching Bond scission | Wear/abrasion Fatigue/fracture Creep/elongation |
| Metals | Corrosion | Wear/abrasion Fatigue/fracture |
| Ceramics | Solubility/dissolution Chemical transformation | Brittle fracture |

Trifluoropropylsiloxane -. Degradation product was carcinogen

In Al2O3, to reduce brittleness of the material, you lower the grain size.

Applications for wanting degradation of material

Tissue engineering

Scaffolds

Langer found a polymer that was easily hydrolyzed in the saline solution found in the body

Found Polyglycolic Acid (PGA), Poly Lactic Acid

Both are native to the body.

Next problem is the degradation rate. Needed to match rate of degradation to the proliferation of the cells.

PGA disappears within 2-4 weeks (begins decomposition immediately)

PLA takes longer to degrade. “Well beyond 4 weeks”

Study the kinetics of cell proliferation and degradation of the scaffold.

Methyl group slows down the decomposition rate – not soluble in water.

Thrombogenicity

Caused failure of jarvik, abiocor, heart valves.

Affect on implant

Wear

Fatigue/fracture

Blood clotting

Valve malfunction

Hindered hydrodynamics

Chemical

- oxidation

- hydrolysis

- adsorption (proteins) – change surface properties

- adsorption – swelling – plasticization (polymers)

- leaching – diffusion

- resorption – biodegrdation

- bond scission

- cross-linking

- change in molecular weight

- change in molecular weight distribution (caused by xlinking/scission)

- degradation products

- biodegradation – enzymatic

- calcification

- change in surface properties

- change in crystallinity

- corrosion

- neutralization of surface charge.

Mechanical

(mechanical degradation)

- fatigue

- crack initiation and propagation

- fracture

- wear

- compliance

(electrica)

- anodic-cathodic reactions

- electrical stimuli (voltage can stimulate bone growth)

- Creep (PTFE) UHMWPE

- Decrease in tensile strength breakage of bonds)

Physical

- change in size

- change in shape

- change in surface topology

- change in optical properties (contact lens- protein adsorption)

- change in hydrophilicity/phobicity

Polymer:

Hydrophilic – changes to the external surface

Hydrophobic – changes to internal area

Promonocyte -> monocyte -> macrophage -> multinuclear giant cell.

The effect of the implant on the body:

(local effects) Wear -> imflammation, macrophages, osteoclast activity (bone necrosis, infection)

(systemic effects) Dr. Patrick Laing first raised the question about the degradation products of co-cr

Transport through the lymphatic system

February 27, 2007

Degradation of products. (Chp16)

Polymers

- PMMA – monomer

  - Inflammation

  - Cardiac arrest

  - Hypertension

Ionic polymerization process – try to get polymerization as complete as possible.

Remove monomers(!)

- PVC

  - Butyltin - Additive causes acid phospatase to collect

  - Plasticizers cause tubing to become softer, allowing for chains to slide over each other

    - Somewhat volatile and can come out of the tubing to contaminated systems

- Silicon gel

  - Lots of physiological interactions with injection

Metals

- Corrosion

  - Cr+n, Co+2, Fe+2,+3, Ti+2

  - Too much breakup can cause tissue poisoning and necrosis

Effects of the implants on the body (no material is inert)

Every implant elicits some physiological response (eg trauma, wound healing)

- Inflammation

- Systemic response

- Protein interaction

- Soft tissue encapsulation

- Hard tissue ingrowth, remodeling

- Change in pH

- Change in electrolyte concentration

- Change in pO2

- Bone resorption

- Sensitivity

- Blood clotting

One form of host response: inflammation

- vasodilation

- increased vascular permeability

- Edema (fluid accumulation

- activation of cellular activity

- initiation of inflammatory response

  - capillary dilation

  - platelet activity

  - coagulation factors

- cellular activity

  - neutrophils/leucocytes

  - macrophages

- remodeling

  - tissue surrounding implant becomes granular

  - collagen activity (often involved in xlinking process)

  - fibroblasts (in soft tissue)

  - osteoblasts (hard tissue)

- capsular formation

  - fibroblasts

  - osteoblasts

Signs for imflammation:

Heat, redness, swilling, pain, loss of function

The role of the inflammatory response:

- isolate/encapsulate

- attack and destroy

any foreign object or device

Cells related to inflammation

- macrophages

  - mononuclear monocytes transform to macrophages (large phagocytes – giant cells)

  - attack and ingrest cellular debris and bacteria

“we consider the macrophage to be the pivotal cell in determining the biocompatibility of implanted materials” – professor james m. Anderson – CWRU

- foreign body giant cells

  - coalesced macrophages

Macrophages – cell receptors

Cytokines: protein-> cells

Secrete chemical products: free radicals, peroxides,

O2., H2O2 – superoxide, OCl-

Secret TGF-B – in the ECM

Activate IL-12

T-cell -> antigen

Role of metal ions in wound healing

Zinc – increases healing rate (positively affects tensile strength of tissue)

Why? May increase collagen cross-linking reaction

May be involved in enzyme reactions

Copper – enhances xlinking of collagen

Partial pressure of O2

Increases with time

At time of wound – 3 torr

Macrophages appear – 10 torr

Fibroblast activity – 20-30 torr

Normal pO2 – 45 torr

pO2 related to collagen synthesis

1 ATM = 760 mm/Hg

Torr = international unit = 1/760 of a standard ATM, dynes/cm^2

pH ranges in the body

blood 7.1-7.4

urine 4.5-6.0

gastric fluids 1.0

intracellular 6.8

interstitial 7.0

pH will:

influence material/host interactions

affect chemical reactions

influence chemical bonds and Van Der Waal forces

Bone response to implants

- metal implants

  - bone plates – block periosteal arterioles

  - intramedullary nails – interfere with medullary arterial circulation

disrupt blood supply – damage blood vessels

revascularization is necessary to maintain bone viability and healing

Bioelectric effect

Wolf’s law

Stress generated potential (SGP)

Bone:

Tensile side is +

Compressive side is –

Piezoelectric

PTFE is also piezoelectric

Piezoelectric theory

Fukuda/yasuda

Streaming potential

Streaming of ions (chlorides, etc) has affect one bone remodeling systems

Electrical stimulation (bone)

- charge stimulates remodeling

+ charge stimulates resorption

Electrical potential drops drastically at site of bone fracture

Polymer implant reactions:

Minimal response:

Silicone rubber, PE, PP, PTFE, PMMA (questionable)

Necrosis:

Some in stitu polymerizing materials (aka PMMA)

Shape and size of implant has effect on body.

Acid phosphatase layer (enzymatic response) is less in a round shape than a rounded triangular shape

Conclusions:

Hard segments are most reactive (-CNO) Thrombogenic

PEO rich surface (soft segment) have low platelet retention

Thrombin adsorption is minimal on PEO segment.

Effects of biomaterials on the body: PMMA monomer

Canine:

Results in rapid dispersion in the body

- leaching of monomer from pmma

- dose 2g/kg body weight

- freshly mixed bone cement

- femoral transcortical plug

- detected monomer level 1 mg/100mg in vena cava in 2 minutes

- peak concentration 3-4 mins, followed by decline

Human:

- peak monomer concentration in 2 mins after implantation

- similar results to canine exp

Infections often associated with prosthetic devices

circulartory shunt – meningitis type infections

ocular prosthesis – conjunctivitis

dental implants - gingivitis

cardiac pacemakers – pocket infections, bacteremia, endocarditis

breast implants – soft tissue infections

joint prostehesis – septic infection

March 5, 2007

Assessment of biocompatibilitly

The testing of biomaterials to determine their safety.

Testing:

- safety

- efficacy

- compliance

Modes of testing:

- in vitro

  - cell culture (2d environment)

  - blood contact tests

  - chemical, mechanical, physical

- in-vivo (animals

  - host – what is the effect of the implant on the body

  - material – what is the effect of the body on the material.

- clinical tests.

  - Functionality – material may be biocompatible but may not be carrying out its intended function.

  - Is the patient reporting pain, illness, blood clotting, etc. does patient actually die? (artificial heart)

- Implant site

  - Subcutaneous

  - Intramuscular

  - Interperitoneal – inside the body cavity

  - Transcortical (through the first layer of bone)

  - Intramedullary

Principles

- determine property of the material

  - material characterization

- biocompatibility tests

  - every time the supplier/procedure is changed, new biocompatibility tests must be done. Must be tied into the quality control system.

- manufacturer – needs an effective quality control system.

Organizations for testing:

- ASTM international

  - ASTM F 04 Materials and medical device committee. (Surgical comm.).

    - Produces test methods

- American association for medical instrumentation (AAMI)

- ISO TC-194 Biocompatiblity committee

- ISO TC-150 Surgical implants.

Maximum implantable dose (MID): maximum amount of implant material (does) that a test animal can tolerate without adverse physical or mechanical effects.

Carcinogenicity test: test to determine the tumorigenic potential of devices, materials, and/or extracts to either a single or multiple exposures over a period of the total life-span of the test animal.

Genotoxicity test: test that applies mammalian or non-mammalian cells, bateria, yeasts, or fungo to determine whether gene mutations, changes in chromosome structure, or other da or gene changes are caused by the test materials, device, etc.

Reproductive and developmental toxicity tests: tests to evaluate the potential effects of devices, materials, and or extracts on reproductive function, embryonic development (teratogenicity) and prenatal and early postnatal development.

Toxic agent: demonstrate an adverse effect on the animal – usually leading to cytotoxicity and cell necrosis.

Cytotoxicity:

- inflammation

  - redness

  - swelling

  - edema

  - pain

  - non-functionality

- invasion of cells

  - leucocytes

  - macrophages

  - lymphocytes

Rating scale:

0 = No visible response

1 = little

2 = some

3 = moderate

4 =

5 = a lot

Classifying toxicity

- acute toxicity

- sub acute

- chronic

Sterilizing with Ethylene oxide (ETO)

- polymers adsorb, upon implantation, the ETO diffuses out and causes inflammation/necrosis

w/ gamma exposure

- changes properties

Processing part (including sterilization) can affect the implant drastically.

F748 – a matrix that tells which road to take in terms of biocompatibility testing.

Testing for:

1. skin irritation (typically using a rabbit) (F719)

Intact skin, abraded (24ours): patch testing.

2. allergic (guinea picg) F720

3. F756 Heymolysis

1. scope –

2. References: astm standards, pharmacopia, fda protocols

3. protocol

Plasma hemoglobin

Knowing the amount of material, can find the hemoglobin index.

Food and Drug Administration

After 1976, medical devices came under regulation from FDA.

510k = provide in vitro/vivo properties of device

PMA = Premarket approval (III) – requires clinical tests of humans

GLP = Good laboratory practice. Documents tell you how to conduct tests, statistical analysis, etc

GMP = Good manufacturing practice

Bjork-Shiley heart valve. Problem was the with the welding of the struts to the ring, they polished over the bad welding.

FDA device classes

Class1: General controls

Not for supporting or sustaining life

Not for preventing impairment to health

No unreasonable risk of iness or injury

EX: bandage

Class2: Performance standards

General controls are insufficient to assure safety and effectiveness

Required to meet applicable standard (section 514)

Class3: premarket approval

Class 1 and 2 controls are insufficient to assure safety and effectiveness

Are life-sustaining or life-supporting

Are implanted in the body

Present unreasonable risk

All class 3 devices are subject to premarket approval (scientific review) requirements

Premarket approval (class 3) (most regulated devices)

Must meet safety and effectiveness requirements

Laboratory studies

- invivo testing for toxicity and biocompatibility in tissue culture

- animal studies: types, numbers, etc

- clinical studies: compliance with IDE (investigational device exemption)

Finally, conduct clinical studies on humans

Evaluate:

- safety and effectiveness

- adverse reactions

- complications

- patient discontinuation

- device failure

- replacements

- analysis of results

- contraindiciations

- precautions

Can standards mitigate or eliminate medical device implant problems?

Silicone gel-filled breast implants

Cyclic fatigue is now being tested

Bjork-shiley heart valves

Manufacture concealed the fact that there were defects in the structure

Total heart replacement devices

Blood clotting is the chief mode of failure

Blood compatibility/clotting tests – would have learned that materials were thrombogenic

Temporomandibular joint implants

Used PTFE, which is bad in compression (creep).

Safe medical devices amendment of 1990

Provides mandatory requirements for medical device tracking

FDA tracking system (proposed?)

Device: Floow device from manufacturer to user

Patient: Lifetime monitoring of device user

How can this be done realistically? What about multidevice patients?

By law, reports must be made if medical devices that contributed to:

death (FDA & manufacturer0

serious injury (manufacturer or FDA if manufacturer is unknown.

Ser

Medical devices subject to tracking

Certain vascular devices (grafts, VAD, pacemakers, heart valves)

Silicone breast implants

Live sustaining devices (ventilators, etc)

Tripartite Agreement (US, Canada, UK) used to regulate devices. Was eventually taken off the market.

Animal tests play a vital role in biocompatibility picture.

Animals in medical research

16-25m animals sacrificed in animals shelters each year in US

Medical research approximate 2% of these numbers

3rd largest number of letters received by congressmen and senators are regarding animal testing. (behind national debt and health care)

Performance test methods (performance standards)

- duplicate body conditions

- realistic test methods

March 6, 2007

International standards for medical devices

ISO TC-194 (technical committee 194)

ISO10993 1-18 evaluation and testing, protocols, etc.

10993-7 “ETO sterilization and residuals”

Identification and quantification of degradation products from polymers/ceramics/metals and alloys

FDA 510k document

Published flowchart that guides certification process

Need to justify that the device is biocompatible

Testing

Standard materials

Certified reference materials

Reference material – not a national standard, “internal reference material”

-Does your material give the same result as the standard reference material

Standard test methods

Scientifically sound

Repeatable/reproducible

- reproducible results through multiple labs

High precision and accuracy (precision vs. accuracy)

Passes interlaboratory testing criteria

Polymer processing

MW -light scattering

-Ultracentrifugation

Mv Viscosity, formula to determine MW

Biocompatibility testing: polymers, ceramics, metals/alloys, composite

Polymers: macromolecule built of many monomer units

Configuration vs conformation

The configuration of the chain refers to the arrangement of the subunits along the backbone of the polymer. Configuration is related to the internal structure of the chain while conformation is used to denote the physical outline or shape of the macromolecule.

Chain folding

Isotactic: all methyl groups (R) are on the same side of the polymer chains

Syndiotactic: methyl groups are on alternate sides of the polymer chain

Atactic: a random distribution of methyl group along the main chain

Amorphous polymers

Description: a mixture of long polymer chains with no particular order

Typical properties: often transparent, poor chemical resistance, softens with temperature, have a glass transition temperature, sensitive to creep

EX: poly styrene, poly nitrile

Crystalline polymers

Description: a mixture of molecules which have ordered or aligned segments along with amorphous segments. Ordered areas are tightly packed

Properties: typically opaque (dense packing), excellent chemical resistance, low friction, amorphous regions soften with temperature, a distinct melt temperature, creep due to amorphous regions

EX PE, PP, fluorocarbons (PTFE), nylon, acetal

Thermoplastic: amorphous and crystalline

Thermoset: cross-linked

Cross linked polymers:

DESC: a single giant molecule interlinked by strong inter connecting bonds

PROPs: typically transparent, excellent chemical resistance, does not soften with temperature., creep does not…

Nondegradable synthetics:

Polyamides, polyesters, polyvinyl chloride, silicones, fluorocarbons, UHMWPE

Biodegradables:

PGA, PLA, etc

Environment changes the surface groups – ie OH or -CH2 groups on the surface

PTFE: fibers

microporous fabric (goretex)

sewing rings for valve struts

Dacron

Sewing ring material

Artificial blood vessels

Delrin

Hemocampatible (76% of implants free of thrombosis)

Poppet wear reported

Adsorption of water

Silastic (silicone)

Early poppets: lipid absorption in starr0edwards valve

Polyurethane

Circulatory assist devices (LvAD)

PDMS coated polymer

Molecular weight are fundamental properties of a polymer sample

Mw =

Mn = weight of molecules / number of molecules

MWD – molecular weight distribuation

Determining molecular weight in a lab

Mn = osmotic measurement. Measure the increase of osmotic pressure across a membrane

Mw =

Properties needed for engineering design of polymers:

Tensile

Strength (ultimate, yield)

Modulus

Creep

Elongcation (ultimate, yield)

Shear strength

Compressive (strength, modulus)

Critical stress intensity factor

Coefficient of friction

Wear characteristics

Glass transition

Fatigue life

Fracture toughness

March 19, 2007

Metals and alloys for surgical implant applications

Properties -> performance

Chemical composition

Mechanical properties

Physical properties

Metal processes – to change the microstructure of the metal.

Casting

A metal or alloy is cast (poured) into a mold. Allowed to cool

Wrought

Plastically deformed metal, shaped by hammering, beating, or pressing

Cold worked at room temp

Hot worked

Forging

A hot working operation to process metals and alloys

Heating and hammering a metal to shape. After foring operations the metal or alloy undergoes an annealing treatment consisting of heating to an optimum temperature and rapidly cooling to meet metallurgical requirements

Recorded history of metals for biomaterials

BC – egyptioans used gold for plates

1829 – Levert, first recorded tolerance study

1930 – Venable discovered Vitallium: Co-Cr-Mo for orthopedic devices

1951 – Leventhal used Titaium for plates and in mesh

1980-1990 Ti alloys with niobium, tungsten

Biocompatibility of Metallic implants is governed by:

Mechanical factors:

Fatigue

Fracture

Elastic modulus

Wear

Chemical factors

Corrosion

Physical factors

Surface characteristics

Physiological factors w/metals

Host response

Chelation – metals become ionic when they corrode, look for free electrons in proteins, enzymes

Sensitivity – Ni+2

Macrophage activation

Vital organs

Local Tissue response

Encapsulation – protein signals fibroblast cell which then deposits membrane around the insult

Granulation – leads to scar tissue

Macrophages – lead to giant cells to deal with objects that are too large for a single macrophage

Necrosis (lysis)

Metals used in implants:

Stainless steels – 316L, 302, 304

Cobalt chrom molybdenum alloys

Titanium and titanium alloys (corrosion resistant, TiO2)

MP35N alloys Co-Cr-Ni-Mo add nickel to allow for easier machining

Nitinol memory metal nickel titanium. Temperature dependent shape properties, stents

Tantalum – used a sa 3d porous structure (scaffold)

Biocompatibility requirements for an implant made of an alloy

- should be corrosion resistant

- implications for biocompatibility

  - Local – necrosis, encapsulation

  - Systemic – heavy metals in system

- should be non-toxic

- suitable mechanical properties

  - wear

  - fatigue

  - fracture

- microstructure – affects mechanical properties

Casting a femoral stem – use lost wax process

Fracture attributable to manufacturing defects:

- inclusions: reduced fatigue strength

- low Mo content (<2%): higher pitting rate

- poor design: sharp corners, holes too close

- excessive porosity: cast 316L SS (need Vacuum casting)

- severe cold working: produced surface cracks

- large grain size: lowered strength

- poor finishing: surface cracks

- improper electropolishing: surface porosity

- improper heat treatment:: lowered strength

- mixed metals (nail/screw: bone plate)

- out of round – femoral heads

- microstructural defects (large grain size)

Cahoon and Paxton, 1968 – 1970 studied hospital purchases, 50% had metallurgical defects

Large grain size becomes “landing strip” for various metals

<.03% C, corrosion resistance of S.S. goes up – develops good Cr2O3 protection/passive layer

>.03% C, form Cr23C6 in the grain boundaries. Depletes amount of Cr to form the passive layer

Characteristics of metals and alloys

Good electrical and thermal conductors

Mobile electrons

Tendency to corrode, oxidize, etc

Comparable tensile and compression strength (within the individual material)

Crystallize readily

More ductile than ceramics

Contain “heavy” metal elements

Examples: S.S. chrome-cobalt alloys, titanium, etc

Energy levels:

Primary Ionic bond 100kcal/mole

Secondary bond  10kcal /mol

Metallic trace elements are essential to life

Vanadium, chromium, manganese, iron, cobalt, nickel, copper, zinc (Atomic numbers 23-30)

When the concentration of the metal ion exceeds the tolerable maximum limit, is it potentially harmful?

Metals and alloys used in surgical implants:

- stainless steel

316L

Iron based alloys with minimum of 12% Cr to improve corrosion resistance

Modulus 28 e6 psi (bone 2-5)

TS 145000 psi

Elongation 10%

There is evidence that S.S. corrodes and metal is released/chelated into surrounding tissue.

Cobalt makes up most of a Co-Cr-Mo alloy ~60%

Haynes-Stellite 21 ASTM F-75

Cast cobalt/S.S. has large grains – weakens materials, but is cheaper

Thermomechanically processed Co-Cr-Mo alloy, far different microstructure

Haynes-Stellite 25 ASTM F-90

Co-Cr-W-Ni alloy

Is Co-Cr alloy a potential carcinogen?

“there is a real problem associated with Co-Cr based alloys.”

“estimates patient with Co-cr implant has 10x the normal chance of developing bone cancer.”

“in animal studies Co-Ccr is associated with the highest cancer rates in an imals.”

“Co-Cr implants have been put in for 50 years. Problems would have surface long before this.”

March 20, 2007

Metal alloys may be carcinogenic in human subjects, but is not totally substantiated.

Titanium and Titanium alloys

1940s introduction as surgical implant

1965 orthopedic implant

Characteristics;

- forms oxide layer

- corrosion resistance

- undergoes wear

Ti alloy is less dense than S.S.

Strength is reasonable.

Types of Ti

Unalloyed titanium

Alloyed ti6al4v

Wrought, annealed

Cast

Forging

Ti6al4w(?) Wrought

4 grades of Ti, each with different amounts of O2. -> changes tensile strength of the alloy

Higher O2 percentage leads to higher tensile strength and lower elongation

Ti 6Al 4V properties

Corrosion resistance

Biocompatibility

Ductility

Fabricablity

High tensile/fatigue strength

Low density

Better modulus match

120,000-150,000 PSI, 8-10% elongation

Ti 6Al 4V

Ti 90%

Al 6%

V 4%

Titanium wear

Generally

- titanium wears more than other metals and causes more wear of UHMWPE

- Nitriding and ion-implantation markedly reduce initial wear but long term wear rates may be much less affected.

Innmunogenicity associated with wear

- Activated t-lyphocytes

- Associated macrophages

- Release of prostaglandin E2 and IL-1 (signs of activated immune system)

Conditions leading to the creation of an electrochemical cell on metallic implants

- two different metals in contact

- variations in O2 concentration

- variations in metal homogenicity

- anodic and cathodic reactions are going on simultaneously

Chemistry of Corrosion

Types of corrosion

- crevice corrosion

- fretting (disruption of any oxide layer that has formed)

- pitting (low O2 concentration), in presence of Cl-

- intergranular

- stress cracking

- galvanic

- cyclic fatigue (goes back to stress cracking)

- generalized corrosion (everything except listed previously)

- erosion – can be chemical or mechanical

pourbaix diagrams – an equilibrium diagram which shows how metals react under conditions of potential and pH.

The nernst equation is used to construct the pourbaix diagram.

Useful for predicting

- direction of reactions

- types of corrosion products

- effect of environment (ph, potential) on surface characteristics

- influence of environmental conditions

metallic corrosion leads to:

- local tenderness

- acute pain

- reddening

- swelling

- chronic inflammation

- changes in cellular metabolism, bone microstructure

- elemental sensitivity

- transport of metal ions

- cell necrosis

Hybrid metals

- drug eluting stents – metals/polymer

- infuse spinal cage

  - cage is Ti cp

  - rhBMP-2 collagen sponge

galling – welds created on asperities due to heat from friction

metal on metal implants, 1980 Sulzer

equal channel angular extrusion (ECAE)

now can take TiCP and match tensile strength of alloy. Can avoid Al, V

Ta has been vapor deposited onto polymers,

Use in acteabular cup can stimulate ingrowth of bone

Hipping

P = 100MPa

T = 1000-1100 C.

Alters microstructure of metal, giving finer grain size, distribution of particles, etc

March 26, 2007

Ceramics

What are ceramics?

Any class of inorganic, nonmetallic products which are subjected to a temperature of 540C and above during the manufacture or use

Includes: metallic oxides, borides, carbides, nitrides, and mixtures of various compounds

Materials which are usually composed of compounds of metallic and non-metallic elements

Al2O3, ZrO2, Carbon (ie, pyrolytic), Glass (ie bioglass), calcium phosphates

Ceramics: Advantages

- largely biocompatible

- low coefficient of friction

- range of reactivity

- dense or porous forms (relating to brittleness, pores serve as crack arrestors)

- coatings over metal

- strong in compression

Disadvantages:

- Brittleness

- Low impact resistance

- High elastic modulus (wolf’s law)

- Weak in tension

- Surface defects (difficult to process and machine without implementing defects

- Low flexural properties.

- Difficult to fabricate

| Metals | ceramics |
| --- | --- |
| Good electrical and thermal conductors Easily lose electrons Tendency to oxidize, corrode, etc Comparable tensile and compression strength Crystallize readily More ductile Contain “heavy” metals | Good dielectrics Accept and share electrons Stable in chmical and thermal environments Stronger in compression than tension Crystallize less readily More brittle Contain ions common in physiological environments (eg Na+, K+, Ca++, Mg++, etc |

Breadth of Ceramics Field

Ceramics can come in two major types

Amorphous – glasses

Crystalline – single and poly crystalline materials

Both types can be composites of sorts

Multiphase glasses

Polycrystalline ceramics, ie pure polycrystalline alumina, refractories

Crystal vs amorphous properties

| Property | Crystalline SiO2 | Amorphous SiO2 |
| --- | --- | --- |
| Tm/Tg | Higher Tm | Lower Tg |
| Strength | Higher | Lower |
| Solubility | Lower | Higher |
| Thermal conductivity | Higher | Lower |
| Hardness | Higher | lower |

Properties of ceramics

Strong inocovalent bonds tend to make ceramics

Strong, hard, brittle, electronically semi conducting through insulation

The ionocovalent bonds can exhibit a wide range of chemical solubilities.

Property development

Properties are developed in different ways

Physical properties are developed by atomic bonding considerations

Microstructural properties are developed by the processing history through ffiring

Part properties are further developed through processing after figing

This is the history that gives us the property seen in an individual device

Processing

Ceramics are processed a number of ways to form the appropriate microstructures and parts -  including

Beneficiating of raw materials

green forming

Sintering

Finishing

Beneficiation

Beneficiating means taking the basic ores or chemicals and making powders or gels that can be used in a ceramic. These methods can include

Fusing of materials (ZrO2)

Chemical precipitation

Sol gel precipitation

Vapor deposition methods

Solution growth

Green Forming

In green forming we take the ceramic stuff and form into the shape we want, processes include:

Cold isostatic pressing (CIP)

Extrusion

Casting (gel, slip)

Sintering, also called “firing” of ceramic ware

Firing is the method by which we make a hard product

This is where the ceramic green powders are coalesced into a single hard piece

Finishing

Finishing is all the post firing stuff needed, in biomedical materials, including

Grinding

Polishing

Joining

Sterilization

Packaging

Microstructural properties

Developed through the processing history of a material

Dependent upon

Grain size, orientation, boundaries

Number of phases

Types of crystals

Grain boundary/amorphous composition

Porosity is a phase

Dopants

Impurity

Alloying agent (metal gets mixed into material, ie ceramic)

Brittleness traceable to the microstructure and bonding:

Ionic bonds with high degree of localization of electrical charges within the lattice

Electrons not movile

Smaller mobility of lattice defects/dislocations (low mobility tolerance)

How to deal with brittleness

control grain size (smaller)

deposit ceramic coating on metal substrate

composite material – add glass/metal fibers to the ceramic matrix. Increase tensile strength and serve as crack arrestors

Al2O3 coatingon 316L SS in water – change in morphology from round to fibrous natures

Wear rates are much lower in ceramic on ceramic articulating surfaces (100x lower on alumina/alumina vs Co-Cr-Mo alloy/UHMWPE)

Zirconias are also popular implant materials

Zirconium Oxide ZrO2

Calcium Zirconate CaO.ZrO2

European concern from alpha and gamma radiation was overcome by testing in US labs

Zirconium Oxide limitations

Wear resistance – poor in comparison to Al2O3

Decrease in strength of material while in a physiological environment, caused by phase transition in the crystal state. tetragonal -. Monoclinic, results in loss of tensile strength

Some of these limitations can be dealt with by stabilizing ZrO2 with Y2O3.

March 27, 2007

Continuing on ceramics

Different types of Al2O3

| Density (g/cm^3) | 3.99 | 3.96 | 3.87 |
| --- | --- | --- | --- |
| Grain size (um) | 15-45 | 1-6 | 5-50 |
| Tensile (MPa) | 206 | 310 | 262 |
| Youngs (GPa) | 393 | 366 |  |

Alumina loses much (50+%) of its adherence strength to 316L SS almost immediately upon immersion in solution.

Specifications of Al2O3

Corrosion resistance (ringers) 0.1 mg/m^2/day

Bone and Al2O3 interface: no fibrous interlayer can be seen at the interface. Good adherence.

| (with stabilization of Y2O3) | Al2O3 | Zr2O |
| --- | --- | --- |
| E Mod (GPa | 380 | 190 |
| Hardness (Mho) | 9 | 6.5 |
| Wear | ++ | + |

Reasons to use Zr2O (used in Europe): wear is not that significant

Hydroxylapatite: Percentage of HA in hard tissue

Bone 60-70%

Cementum 70%

Dentin 77%

Enamel 98%

- Wide variation in the mech. Properties of HA (different manufactures, grades, etc)

|  | Jarcho | Cato |
| --- | --- | --- |
| Compressive strength (mpa) | 196 | 294 |

Ti6Al4V – immediately after placement, Titania(s) is formed

Titanium oxide TiO2, TiO, Ti2O3

Calcium Titanate 3CaO.2TiO2

Isotropic carbons used in clinical devices

Pryolytic carbon (low temperature isotropic, LTI) aka Pyrolite carbon

Glassy carbon (vitreous carbon, polymeric carbon)

Vapor-deposited carbon (ULTI carbon) aka Biolite carbon

Vapor deposition – typically deposit the carbon onto polymer/ceramic substrate

Carbon fibers and composites

Issues: characterization, strength, fracture toughness

ULTI carbon is strongest in tension among glassy, LTI, LTI w/Si carbons

Bioglass

~45 % SiO2, 6% P2O5, 24.5% CaO, 24.5% Na2O

Why P2O5? -> calcium hydroxylapatite contains P.

By reducing CaO and adding CaF2 (12.25%), solubility was drastically changed

A triangular phase diagram (SiO2, Na2O, CaO) was created to determine bone bonding results. Good results were found at somewhat equivalent amounts of the three materials (see composition above)

Mechanism of bonding between bone and glass ceramics

Thre is a gap between bone tissue and the glass-ceramics immediately after operation. First the surface of the glass-ceramics becomes irregular because of dissolution between 5 and 10 days after implantation. Second a Ca-P layer is formed on the irregular surface of the glass-ceramics. The surface of the Ca-P layer at near the bone tissue is smooth. Bone tissue grows towards the Ca-P layer between 10 and 30 days after implantation.

Strong bonding adherence of collagen to bioglass surface.

Collagen

Collagen represents about 25% human tissue

15A wide, 400A periodicity of banded patterns.

Importance of TiO on top of Ti alloy

Formation of Apatite on Ti and TiO2

The oxide layer formed on Ti implants in the body increases in thickness and attracts minerals from the surrounding physiological solution.

Proteins are first adsorbed on TiO2, then mineral ions diffuse through the adsorbed protein layer.

Plasma Spraying – most of the coatings that are made from ceramics are made by some high energy phenomenon.

April 2, 2007

Composites: materials composed of two or more different constituents, each of which contributes specific property characteristics that enhance the performance of the product.

Composites enable the design of surgical implants with mechanical and physical properties that are not attainable using single materials alone.

Other advantages of composites:

- Mechanical properties may be varied

  - Lower modulus of elasticity than metals

  - High tensile, flexural, and fatigue strength

- Improved biocompatibility

- Reduced corrosion

Types:

Polymer – fiber, filler

Bioglass – stainless steel

Al2O3 and TiO2 – metal (plasma sprayed)

Hydroxyapatite/fuorapatie

Bisphenol-a/glycidyl methacrylate (BIS-GMA)

SiC/C - carbide

Carbon Fiber/reinforced C

Matrix:

Polymers

Thermoset – epoxy, phenolic

Thermoplastic – PMMA, PE, PP, Polysulphone, Polyester, Kevlar

Metals

Al, Ti, Ni

Inorganics

Carbon, Al2O3

Combinations

Polymers/metals/inorganics

Metals/organics (coatings)

M/M

Matrix Factors

Modulus

Transfers load to fibers

Supports Fibers, maintains their position

Al2O3, ZrO2, and TiO2 coatings on metal implants

Reduces corrosion and M release

Reduces M sensitivity and wear

Applications

Skeletal systems

Fixation devices

Total hip endoprotheses

Bone cement

Total knee

Etc

Dental Implants

Resorbable Sutures – PLA/PGA

Artificial skin – mixture of nylon and silicone

Characteristics:

Less rigid than metals

High fatigue strength

Promotes bone healing/remodeling

Pain reduction (lower modulus)

Good-excellent biocompatibility potential (aramid fiber/pmma)

Wide latitude in design (shape, varying stiffness, fiber volume, etc)

Chemically resistant (ie C-fiber/polysulfone)

Eliminates metallic corrosion

Fiber-reinforced composite characteristics

Improved strength

Fatigue resistance

Reduced crack propagation

Adjustable elastic modulus

High strength to weight ratio

Stress transfer from matrix to fibers

Composite types

Particulates – isotropic

Fibers – isotropic or anisotropic

Laminates – anisotropic

Fiber factors

Length

Diameter

Orientation

Quantity (%fiber content)

Fiber properties – surface properties

Aspect ratio l/d  - 3:1 ratio is a carcinogen (asbestos). Don’t want these fibers to be released

Mineral fibers

Natural

Zeolites - mineral

Asbestos

Synthetic

Bioglass

Ceramic

Carbon

Organic

Natural

Silk

Cotton

Synthetic polymer

Metallic fibers

SS

Ti 6/4

Co-Cr-Mo

Carbon fiber reinforced bone cement

Advantages

Improved strength and fatigue life

Improved modulus of elasticity

High flexural strength

Reduced exotherm

Reduced thermal coefficient of expansion

Disadvantages

Loose fibers

Bonding difficulty

PMMA

Exothermal reaction – add a filler

Improve catalyst (benzoyl peroxide)

Improve mech properties by selecting fibers/fillers

Kevlar, 5x TS of SS by weight

In contact with tissue, Kevlar forms a thin fibrous layer between polymer and tissue

Coated metal composites

Bioglass +316L SS

Metal – bioglass interface is critical to performance of implant

Pretrement of metal surface assues proper bond to bioglass (cleaning, deoiling) 3% HF

Al2O3 and TiO2 coatings on metal implants (plasma spray technique)

Reduces metallic corrosion and release of metal ions

Minimizes imflammatory response

Metal ion toxicity can be reduced (Ni, Cr)

Carbon fiber randomly reinforced UHMWPE

Chapped carbon fibers, wt%: 2-40

UHMWPE: balance

CF/90% PE are in clinical use.

April 3, 2007

More composites

Fiber directions

Composite factors

Fibers must be firmly bonded to matrix

Glass fibers – silane coupling agent

Carbon fibers – organic coatings

Coefficients of thermal expansion must be similar

Interface factors

Nature of matrix/fiber bond

Integrity of bond

Usually the medical device composites consist of high strength fibers (polymers, metals or ceramics) in a ductile matrix (ie polymers) some fibers may also be incorporated in ceramics

Composite femoral stem – SS fibers with a polysulfone matrix

Fibers of the matrix separated under stress and was never successful (did not match elastic mod of bone)

Limitations hindering further development of composites as surgical implants

Problems dealing with adhesion of fibers/matrix

Polymers – hydrophilic will swell (leading back to adhesion problems)

Complex manufacturing procedures

Not widely accepted for use in implants.

Scarcity of standards for testing – inspecting and testing adhesive bonds of surfaces.

Polypropylene matrix with untreated glass fibers:

Fibers pull cleanly out of resin

With silane coating – failure is solved

Silane has =CH2 coupling site to polymer matrix

Other factors leading to good bonding between polymer and filament

Low contact angle between polymer and fiber

Hydrophilic surface has lower contact angle than hydrophobic (spreads out more)

Want low viscosity resin

Needs to be clean and free of dust.

Avoid surfaces with (micro)cracks and imperfections

Want moderate roughness to the surface/fiber interface -> greater surface area

Coefficient of friction between filler and matrix should be the same

Tests on Molded composites

Uniform dispersion of randomly oriented carbon fibers

Fiber size to be determined by composite manufacturer

Mechanical properties

UTS

UYS

Ultimate elongation

IZOD impact strength (ramrod attached to set of rates, collides into plastic material)

Biocompatibility in accordance with ASTM F-748

Cell culture

Intramuscular implantation

Long term implant

(no site specific testing)

Biocompatibility of different classes of composites

Class I – minimal biological response

CF/reinforced CF

CF/Polysulfone (biologically minimal response, but not mechanically)

CF/PMMA

Class II – Active tissue response

Bioglass SS fibers

CA Phosphate/UHMWPE

Class III – Bioresorbable response

PLA/PGA

CaPhos/PLA

AcrlyNitrile – thermal decomposition for making graphitized structure (elemental carbon)

Pyrolytic carbon formation

Dr. David Hungerford

Dr. Fromdoza

Interested in repairing femur with metal rod into intermedullary canal of femur

Wanted composite material to match elastic mod of bone, and be biocompatible

90% PEEK polymer

10% glass fibers

Took MG63 osteoblasts to evaluate biocompatibility.

Osteocalcin is a marker of osteoblast activity

April 9, 2007

Chimeric Neomorphogenesis

The process of engineering a tissue or organ in situ by placing dissociated cells onto synthetic biodegradable scaffolds and placement in a host to permit growth, function, and vascularization.

Tissue engineering: an inter disciplinary field that applies principles of engineering and the life sciences to the development of biological substitutes that restore, maintain and improve the function of damaged tissues and organs.

Goal of tissue engineering:

Treat disease or malfunctioning body parts by transplanting specific cells and tissues that have been engineered in the laboratory.

Tissue engineering

-design

-specification

-fabrication

8-10 million transplants performed

40-90 million hospital days

Liver:

30,000 patients/year need liver attention

3000 liver donations/year.

750k new diabetic patients are identified each year.

150k die from diabetes.

1m die from cardiovascular disease/year

Burn patients

Current treatment

- autograft

- allograft (cadaver skin)

  - antigenic

  - potential disease carrier

  - limited supply

Tissue engineering:

- derma graft

  - fibroblast cells (secrete proteins and growth factors)

  - nylon mesh

  - silicone membrane

Tissue engineered implant:

A biologic-biomaterial system designed to restore, modify or augment tissue or organ functions

Standards development:

ASTM F04 – medical and surgical materials and devices Division IV: TEMPs

Focus:

Biological components (cell, tissue, cellular, product and/or biomolecule

Biomaterials components (natural/synthetic)

Preclinical/clinical assessments

Nomenclature

Limitations of current state of the art valves:

Mechanical valves

Foreign body response

Lack of growth

Mechanical failure

Need for lifelong anticoagulation

Thrombosis

Tissue valve (xenograft)

Foreign body response

Lack of growth

Short durability

Calcification

Allogenic valve (homograft)

Foreign body response

Lack of growth

Donor organ scarcity

Rejection

Porcine heart valve (aortic position)

Cryolife, Inc. Atlanta GA: syner graft procine heart valve

Dr. Mark O’Brien performed first implants of tissue-engineered heart valves

Two women patients in Brisbane, Australia

Technology

Porcine heart valve depopulated of cells

On implantation, will repopulate with patients heart cells

Allows trans-species transplant without use of immunosuppression

Tissue engineering system:

Cells (millions

Scaffolds

Bioreactors

Culture vessels

Cell and tissue sourcing (problems)

Autologous – aseptic harvesting, rigorous record keeping (time limits!)

Allogeneic – aseptic harvesting, rigorous record keeping, adventitious agents

Xenogeneic – aseptic harvesting, rigorous record keeping, adventitious agents, retroviruses.

Functional evaluation of tissue engineered cells/tissues

Cell metabolism – glucose consumption, lactose production

Cell damage – concentration of lactate dehydrogenase

Extracellular matrix – synthesis rates 9radio labeled)

Measure mechanical properties (eg cartilage)

Electrophysiological properties (eg cardiac tissue)

Scaffold requirements:

Provide structural framework for selected cells

Biocompatible

Resorbable

Promote cell differentiation, function, and growth

High porosity

Allow exchange of nutrients and waste products

Prevent entry of immunoglobulins and lymphocytes of immune system

Scaffold design

- surface chemistry

- surface charge

- surface morphology

- resorbability

- cytotoxicity

examples of resorbable polymer scaffolds

polyesters

- PGA, PLA

- PGA/PLA co-polymers

- Poly hydroxyl butryl acid (PHB)

- Poly hydroxyl valeric acid (PHV)

- PHB/PHV copolymers

- Poly (e-caprolactone)

April 10, 2007

Molecular weight Mv = molecular weight determined by viscosity

Mn = number average molecular weight

Tensile strength decreases with decreasing MW

MW decreases with time in vivo/vitro.

Importance of resorbability

- eliminates chronic foreign body reaction

- reduces chance of infection

- leads to restoration of normal tissue

Crystallinity and thermal properties tof PGA PLA and copolymers

|  | % xtalinity | Tm | Tg |
| --- | --- | --- | --- |
| PGA | 46-52 | 225 | 36 |
| 90:10 PGLA | 40 | 210 | 37 |
| 50:50 PGLA | 0 | None | 55 |
| PLA | 37 | 185 | 57 |
| dl-PLA | 0 | None | N/A |

Further addition of LA leads to totally amorphous material. Much more soluble than crystalline material. Gives a time window to match up cell proliferation rates with degradation of matrix.

Scaffold surface (pore size) important features

- % porosity (usually between 70-80%)

- Pore shape (round, square, etc)

- Surface area/properties

Pore size

5um – neovascularization , capillaries.

5-15 um – fibroblast growth

20-125um – epithelial cell growth (tissue engineered skin)

100-300 um – osteoblast growth

Cell Adhesion – a key issue in tissue engineering

- Polymer surfaces (scaffolds) generally do not have the ability to interact with cell-surface recptors to affect cell adhesion

- When implanted polymers adsorb adhesion proteins they support cell adhesions

- Cells bind (adhesion) to adsorbed proteins rather than the polymer

Intrigens – surface receptors, peptide subunits, fibronectin.

Carbohydrates and cell adhesion

- mono saccharides are incorporated into polymers to promote cell adhesion

  - lactose bond to polystyrene binds mammalian hepatocytes

- carbohydrates bind to lectins (cell adhesion proteins)

  - lectin UEA-1 has been bound to polyethylene terephthalate

    - increased endothelial cell attachment 100 fold

    - fibroblasts, monocyte, smooth muscle adhesion was reduced

Oligopeptides – do not denature when they get to the surface

Covalently bonded to scaffold

Cell surface receptors

Oligopeptide sequences used for cell adhesion

RGD – fibronectin, collagen, laminin

Integrins:

Composed of alpha and beta transmembrane subunits

Selected from 16 alpha and 8 beta subunits

Produce more than 20 different receptors

Integrinn mediated adhesive reactions are involved in many cellular functions

Leukocyte homing and activation

Hemo statis (blood clotting)

Bone resorption

Response of cells to mechanical stress

Tumor growth and metastasis

Programmed cell death

Biological signals

Cell signaling and cellular responses

Signal transmission by surfaces, bioactive agents, ligands

Mechanisms

Biomaterial encapsulation

Phagocyte response

Bone deposition

Failure of tissue engineered implants, eg immune response

What is a growth factor?

A powerful regulator of biological function

A biologically active protein that promotes cell growth…

Polypeptide growth factors

Increases cell proliferation, migration, aggregation, and differentiation

Bind to cell-surface receptors to stimulate cellular activity

FGF fibroblast growth factors, EGF epidermal growth factors, PDGF platelet derived growth factor

Tissue engineering bioreactors

In vitro cell culture systems

Promote cell growth on 3d scaffolds

Provided needed nutrients and gases

Efficient mass transfer to cells/tisssues (flow and mixing)

Subject developing cultures needed to physically stimulate

Static and rotating bioreactors

Bioreactors: Static/rotating

Cartilage and cardiac tissues functionally superior in rotating bioreactors vs static reactors

Hydrodynamic forces affect cells via

Stress fields, stretching, pressure fluctuations.

Smooth muscle

Seeded tubular silicone scaffolds

Built with Pulsatile radial stress

Produced arteries (<6mm diam)

High rupture strength

High suture retention

Bioreactors with mechanical stress components

Cartilage

Dynamic compression at physiological frequencies

- promotes extraceullular matrix synthesis

- increases glycoaminoglycan synthesis (GAG)

- repaired bone defects in rabbits

Cardiac tissue

- increased size and function (contraction)

- produced growth hormone

Tissue engineering of the liver – hepatocytes

Seeded on polymer scaffold (5e7 cells/ml)

Polymer scaffold: PGA PLA PGLA PVA

Scaffold implanted in host

Hepatocyte growth factor (HGF)

- increased cels up to 30x in 4 weeks (rats)

- sustained release of HGF

Islet cells for diabetic patients

Desmos, Inc. San Diego

Produces pancreatic islets

Uses ECM protein laminin-5 as growth substrate

Microislet, La Jolla

- Provides encapsulation technology

- Alginate scaffold

Osteotech

-provided HA and Ti plasma spray coatings for orthopedic and dental implants

Stryker biotech

- markets a human recombinant osteogenic protein for nonunion fractures of long bone

Tissue regeneration

- isolates mesenchymal cells of bone marrow, grows them on a scaffold in a bioreactor obtaining a human ligament product

Bovine – derived cultured skin

Ortec International

Porous matrix of bovine collagen

Seeded with two layers of human dermal and epiththelial cells

Strategies:

- replace only those specific cells needed for function

- provide tissue inducing substances eg growth factors, signal molecules, etc

- ccells incorporated in a matrix; synthetic or modified natural materials.

  - Open cell matrix: implanted in body, become integrated with host tissue

- Closed matrix: cells are isolated from the body by permeable membrane

  - Permits nutrients, gas exchange to enter

  - Prevents entry of antibodies, immune cells

April 16, 2007

Skin - organogenesis inc Apligraf

Cartilage carticel genzyme

Encapsulated islet cells, hepatocytes, urothelial cells

Tissue engineering of bone – why is it important?

- millions of surgical procedures, bone defects, non union

- autogenous bone grafting

- allogenic bone grafting

Scaffold: designed to be osteoconductive – addition of growth factors, etc

Supports angiogeneisis

- promotes cell proliferation, differentiation

Distraction osetogenesis – application of dynamic force to cell colony

- Pulsaltine

- Cyclic mechanical testing

- longitudinal force

Mesenchymal stem cells (MSC) (pluripotent)

Osetoprogenitor cells

- more advanced and can only produce bone

- low population in bone marrow

Growth factors are specific proteins secreted by cells and are signaling molecules that interact with target cell receptors.

Bone Morphogenic factor

- rhBMF 1-15 have been identified

- BMF-2,3,7 involved in bone repair and stimulating MSC differtiation into bone

- BMP-3,4,7,8 active during day 14-21

  - Ostoblastic recruitment

  - Resorption calcified cartilage BMP -5,6-active during day 3-21

Vascular endothelial growth factor (VEGF)

- stimulates vascularization

- promotes new bone formation

transforming growth factor Beta (TGF-B)

- stimulates MSC proliferation

- stimulates bone regeneration

- increased mechanical properties of new bone

Roles of growth factors (proteins)

- cell adhesion

- cell proliferation

- cell migration

- cell differentiation

Tissue engineering factors:

- cells

- scaffolds

- bioreactor

- upscaling (from “test tube” to large volume)

- preservation

- imuno-compatibility

Osteogenic – bone syn cells

Osteoconducdtive

Stem Cells:

1998

- James Thompson, University of Wisconsin established the first human embryonic stem cell line.

- US patent issued

Legal issues

- Wisconsin alumni research foundation (WARF) licensed Geron Inc to commercialize stem cell by-products: liver, muscle, nerve, pancreas, blood and bone

- WARF is suing geron, Inc. contesting legality

Ethical issues

- is it morally ethical to use human embryonic stem cells for research and applications

- Dr. Roger Pedersen UCSF is moving his laboratory to Cambridge university, enging in a more “receptive environment”

- Patient’s right to know (patient autonomy)

- Manufacturer’s obligation to patient

- FDA’s oblicgation to protect and inform patients

- Engineer’s obligation to

  - Company

  - Patient

  - Engineers code of ethics

Other issues

- the provider’s demand for any new inventions

- etc

What is needed?

Standards: reference biomaterials and test methods to improve precision and accuracy of laboratory test results

What is a reference material?

A well characterized material that by a standard test has been determined to elicit a reproducible, quantifiable host or material response.

What is a standard test?

A test based on established scientific and engineering princibples that has been demonstrated to yield accurate and reproducible results

Challenges for tissue engineering

Regulatory

Are products safe and effective?

Mechanism for evaluationg “combination” products

Health insurance coverage

Are these products significantly better than current treatments?

Are they affordable?

Manufacturing and quality control

Source of materials (animals, variability of natural products, recombinant products).

Complexity of material preparation. Scale up stem cell culture.

Short shelf life

Fundamental understanding

Cell response to environmental stimuli: imperfectly understood.

Cell types:

- fibroblasts

- myoblasts

- endothelial cells

- valvular interstitial cells

Scaffolds:

PGA (resorption time 2-4 weeks)

PLA (18-36 months)

PLGA (8-12 weeks)

PHA PolyHydroxyAlkanoates – ex Poly-4-HydroxyButyrate

PGA-PHHB

Natural

Collagen

Alginate

HA

BioReactors:

Want dynamic feature –

Some key points to improve performance and minimize failures:

- rigorous design criteria and testing

- meaningful materials specification requirements

- standard test methods

- quality control system

we need statistical data:

- how many implants of a particular type are in patients?

- How many implants are failing?

- What is the mean time to failure

- What is the principle mode of failure?

Need to develop a national and internation network of cooperating surgical implant retrieval and analysis groups to provide validated data on device failures and cause of failures

The future:

Biomaterials and biotechnology espanding markets due to:

- graying population

- monderization of developing nations health care systems

- new tissue engineered product lines

- increased use of drug delivery systems

- research and development will lead to new innovations

Bioaterials and biotechnology

- greater involvement of molecular biology, biochemistry, biomimetics, nanobiomaterials and drug delivery systems will lead to remarkable advances

- growth factors, cell signaling, stem cells, biocompatibility, improved synthetic materials, bioinformatics

April 23, 2007

Biocompatibility

Material failures

- Surgeon

- Patient

- Processing

Breast implants – reports of adverse effects

- fibrous encapsulation

- capsular contracture

- interference with mammograms

- silicone bleed

- capsule rupture

- formation of calcium deposits

- calcification

- implant shifting position

- silica filler

- arthritis-like symptoms

- corrective surgery

- polyurethane degradation

Countermeasures to prevent capsule formation

- drains for hematoma and fluid accumulation

- antibiotics

- capsulotomy (exercise)

- retromuscular positioning

- steroids

- polyurethane coating

Silicone breast implants? Do they cause connective disease? Immune related disorders? Cancer?

“There is currently no evidence regarding silicone toxicity”

PTFE, UHMWPE

PMMA bone cement (still in use)

Porous surfaces – bypass PMMA depends on bone integration for mechanical stabilization

Metals 316L SS, Co-Cr-Mo, Ti6Al4V

Ceramics Al2O3 ZrO2

How can we overcome some of the wear problems with UHMWPE?

- eliminate loosening of femoral stem

- modify surface of acetabular cup

- eliminate UHMWPE degradation via gamm sterilization

- control Mwd of UHMWPE

- strict quality control of processing conditions for production of UHMWPE

- modify design of acetabular devices

- improve surgical implantation technique

- use another material

affinity index – length of bone in contact with surface/length of implant surface

Delrin

Affect on implant

Wear, fatigue/fracture, blood clotting, changing hydrodynamic flow of blood

Thrombus formation

Adsorption of plasma proteins -> adhesion of platelets -> aggregation of platelets -> fibrin formation -> thrombus

Albumin

Adsorption decreases platelet adhesion

Decrease thrombogenicity, blocks fibrinogen adsorption which attracts platelets

Albumin does not contain RGD sequence, and amino acid sequence common to “adhesive proteins”

Albumin must occupy most of the surface (>98%) to be anti-thrombogenic

Intrinsic pathway – XII Hageman factor

Extrinsic – tissue factor, VII

Ti6Al4V

Corrosion resistance, biocompatibility, ductility, fabricability, high tensile, fatigue strength, better modulus match to bone

Polymer additives – plasticizers, antioxidants, fillers, cross-linking agents, lubricants, mold release agents, etc